Dental Clinical Audit

Decontamination in Primary Care

Dental Practice

(only one dentist per practice to complete this audit per year, in conjunction with cross-infection lead)

Audit start date:

Completion date:

April 2015

1. Decontamination in Primary Care Dental Practice

Patients deserve to be treated in a safe and clean environment with consistent standards of care every time they receive treatment. It is essential that the risk of person-to-person transmission of infections be minimised as much as possible.

The Department of Health has produced guidance (a Health Technical Memorandum HTM 01-05) on decontamination in dental practice. It has been produced after wide consultation and reflects a commitment to improving standards in dental practices.

By building on existing good practice – this guidance can help us to deliver the standard of decontamination that our patients have a right to expect.

The policy and guidance provided in this Health Technical Memorandum are aimed at establishing a programme of continuous improvement in decontamination performance at a local level. The guidance suggests options to dental practices within which choices may be made and a simple progressive improvement programme established. It is expected that by the end of the first year of the implementation of this guidance, all primary care dental practices will be working at or above the essential quality requirements described in this guidance.

The Care Quality Commission (CQC) has regulatory responsibility to ensure that the requirements for registration are met. This includes the provision of a safe, clean environment and appropriate decontamination of dental equipment.

Registration places a strong emphasis on effective management and self-audit. These criteria are part of clinical governance and are defined in the ‘Health and Social Care Act 2008: Code of Practice for the NHS on the prevention and control of healthcare associated infections and related guidance’ (the HCAI Code of Practice). This Health Technical Memorandum is specifically referred to in the HCAI Code of Practice.

IT IS STRONGLY ADVISED THAT YOU OBTAIN A COPY OF THE HTM 01-05 GUIDANCE AS IT WILL HAVE CONSEQUENCES ON THE WAY DECONTAMINATION IS UNDERTAKEN AND WILL PRESENT SIGNIFICANT PROBLEMS TO SOME DENTAL PRACTICES  IN MOVING TOWARDS “BEST PRACTICE”

All practitioners will have received a hard copy of the guidance eventually but in the meantime it can be downloaded from the internet by typing or copying and pasting the following link:

http://www.dh.gov.uk/dr_consum_dh/groups/dh_digitalassets/documents/digitalasset/dh_097679.pdf

The guidance will help you to complete the audit.

A SUMMARY OF THE GUIDELINES

To help dental practices to improve their decontamination procedures, the guidance introduces specific benchmarks by which compliance with

- essential quality requirements

and

- best practice

can be achieved and demonstrated

EVERY PRACTICE SHOULD BE CAPABLE OF MEETING THE

ESSENTIAL QUALITY REQUIREMENTS,

THAT IS:

- Regardless of the technology used, the cleaned instruments, prior to sterilization, should be free of visible contaminants when inspected. Instruments should be reprocessed using a validated decontamination cycle including: cleaning/washing (in terms of manual cleaning, this includes having a written protocol, a validated steam sterilizer, and at the end of the reprocessing cycle they should be in a sterilized state.

- Reprocessed dental instruments should be stored in such a way as to ensure restraint of microbiological re-colonisation. These measures should be backed by careful controls on the storage times to which instruments that are less frequently used are subject.

- Practices should audit their decontamination processes quarterly using an audit tool .

- Practices should have in place a detailed plan on how the provision of decontamination 	services will move towards best practice.

(An expanded list of essential quality requirements is given in paragraph 2.6. of the    	guidance).

TO DEMONSTRATE BEST PRACTICE,

FURTHER IMPROVEMENTS ARE REQUIRED IN THREE MAIN AREAS:

- A cleaning process that should be carried out using a validated automated washer-disinfector.

- The environment in which decontamination is carried out should be such as to minimise the risk of recontamination of instruments and the possibility of generating aerosols, which may reach patients or unprotected staff. For best practice, the decontamination facilities should be clearly separate from the clinical treatment area. This implies the use of a separate room or rooms for the accommodation of clean (output) and dirty (input) work. In these facilities, the room(s) should be used for this purpose only and access should be restricted to those staff performing decontamination duties. However, plant and equipment not necessarily used for decontamination may be located in these rooms (but preferably in the dirty room) provided it can reasonably be shown that the devices do not conflict with the requirement for a clean environment.

- The storage of reprocessed dental instruments in a simple but carefully designed facility clearly separate from the clinical treatment area is an important best practice improvement. The facility should take account of the need to reduce re-colonisation of sterilized instruments and also make the identification/selection of instruments easy. This storage facility will ordinarily be part of the clean area within the decontamination room(s).

Not all practices will, at present, be in a position to adopt best practice recommendations. However, every practice will need to assess the improvements they need to undertake to move towards these and prepare a plan to implement the changes.

While all practices should now have achieved essential quality requirements, no schedule for attainment of best practice is provided in this guidance for the present. Findings from the 2009 National Survey of Dental Decontamination will be used to guide the Department of Health (DH) and the dental community in terms of reasonable timescales. It is recognised that not only are improvements in premises and equipment required to achieve higher standards, but also changes in practice management and the culture in which patients are treated by the dental team are necessary.

The guidance is based upon a principle of continuous improvement in the quality of decontamination practices and the environment used. Where dental practices use the same room for patient treatment and decontamination (essential quality requirements), they need to develop a plan that facilitates a move towards a separate and controlled decontamination room. This plan will normally also contain statements on staff training and development to suit work in a dedicated decontamination room or suite.

In addition, the plan should realistically outline the way forward in relation to best practice requirements, for example:

- measures to purchase and incorporate a washer/disinfector;

- the separation of decontamination processes from the patient treatment area.

As you will appreciate these changes to the decontamination process will require changes not only in the infrastructure but the management and staffing of dental practices. This audit is the same audit tool produced by the Dept of health to accompany the guidance. Practices are expected to do a self assessment audit using this tool so make sure you keep a copy to avoid repeating the exercise. The idea of the audit is NOT to answer all the questions correctly but to answer them truthfully in order to create an action plan firstly to ensure that “essential quality requirements” are met and secondly to start planning how your practice will proceed to “Best Quality”

The guidelines are significantly more comprehensive than before and as a result many practices will identify areas of weakness that need attention. The purpose of the audit is to help you plan how to raise your standard of decontamination up to the required level

The Care Quality Commission will be looking to ensure that dental practices both NHS, mixed and entirely private meet these requirements.

Practitioners worried about the effects of these requirements on their practice should contact the BDA or their Local LDC for advice.

1. Decontamination in Primary Care Dental Practice

AIMS AND OBJECTIVES:

Aims:

- To enable dental performers to evaluate their standard of cross infection control against the new guidelines

- To use the results to promote discussion and decide on any necessary changes to practice and procedures

Objectives:

- Any necessary changes to achieve essential quality will be planned and implementation started

- An action plan to move to best practice will produced

Audit on Decontamination in Primary Care Dental Practice

METHOD:

Complete the self assessment audit referring to the HTM guidance where necessary

When you have completed the audit identify those areas that require attention to achieve the “Essential quality requirements”

Produce an action plan with dates for implementing any necessary changes

Identify those areas that need attention in order for your practice to move towards “best practice”

Not all practices will, at present, be in a position to adopt best practice recommendations. However, every practice will need to assess the improvements they need to undertake to move towards these and prepare a plan to implement the changes.

Prepare a plan which should realistically outline the way forward in relation to best practice requirements, for example:

- measures to purchase and incorporate a washer/disinfector;

- the separation of decontamination processes from the patient treatment area.

1. Decontamination in Primary Care Dental Practice

| Decontamination in Primary Dental Care Clinical Audit Feedback: |
| --- |
| Were the following AIMS & OBJECTIVES ACHIEVED | Yes | No |
| To enable dental performers to evaluate their standard of cross infection control AGAINST THE NEW GUIDELINES |  |  |
| To use the results to promote discussion and DECIDE ON ANY necessary changeS TO PRACTICE AND PROCEDURES |  |  |
| aNY NECESSARY CHANGES TO ACHieVE ESSENTIAL QUALITY WILL BE PLANNED AND IMPLEMENTATION STARTED |  |  |
| AN ACTION PLAN TO MOVE TO BEST PRACTICE WILL BE PRODUCED |  |  |

Action Plan and Changes made after Audit: “Essential quality requirements:” Action required Action Plan with dates “Best Practice” Actions required with timescales

How useful did you find this Dental Clinical Audit? Please tick one of the following: No use Useful Very Useful

Any comments on this Structured Dental Clinical Audit especially if you ticked no use:

- Prevention of Blood borne virus exposure						  CAP Ref:

The risk of blood borne virus exposure (including Needlestick injuries, bites, splashes involving blood or other body

fluids) is managed to prevent infection

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have a policy and procedure/s in place for the prevention and management of blood borne virus exposure, including management of spillages, sharps and inoculation incidents in accordance with national guidance? |  |  |  | 2.6 |
| Have all staff received training in relation to the prevention and management of blood borne virus exposure? |  |  |  | 1.25 9.1 9.5 |
| Have all staff at risk from sharps injuries received an occupational health check in relation to risk reduction in blood-borne virus transmission and general infection? |  |  |  | 2.6 |
| Can decontamination and clinical staff demonstrate current immunisation with the hepatitis B vaccine e.g. documentation? |  |  |  | 2.4s 8.8 |
| Is sodium hypochlorite available for blood/bodily fluid spillages and used as per manufacturers instructions? |  |  |  | 6.74 |
| Are sharps containers correctly assembled? |  |  |  |  |
| Are sharps containers in use labelled with date, locality and signed? |  |  |  |  |
| Are sharps containers filled beyond the indicator mark? |  |  |  |  |
| Are sharps containers locked with the integral lock when filled to the indicator mark? |  |  |  |  |
| Are full sharps containers stored in a secure facility away from public access? |  |  |  | Appendix 1 |
| Is there a readily accessible protocol in place that ensures staff are dealt with in accordance with national guidance in the event of blood borne virus exposure? |  |  |  | 2.6 |
| Are inoculation injuries recorded? |  |  |  |  |
| Are disposable needles and disposable syringes discarded as a single unit? |  |  |  |  |
| Are sharps containers available at the point of use and positioned safely (e.g. wall mounted)? |  |  |  |  |

- Decontamination

Medical devices are decontaminated prior to use and any associated risks are safely managed

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have a policy or procedure that includes all appropriate aspects of decontamination within the practice e.g. cleaning, disinfection, inspection, packaging, disposal, sterilization, transport and storage of re-usable and single use instruments? |  |  |  | 2.6 |
| Have all relevant staff received training which they are expected to perform for decontamination including correct use of equipment? |  |  |  | 1.25, 2.4q 3.7, 3.16 |
| Is a record of any instruments that cannot be reprocessed in accordance with your local decontamination policy? |  |  |  |  |
| Are all wrapped, sterilized instruments dated with the use-by date? |  |  |  | 1.23, 1.9 4.24, 4.28 |
| Does the practice have a nominated lead responsible for infection control and decontamination? |  |  |  | 2.4c 9.3 |
| Has the registered manager a written statement of duties with specific reference to equipment validation? |  |  |  | 11.5 |
| Is there a procedure for transportation of instruments to and from other locations, that ensures the segregation of contaminated instruments from clean/sterilised instruments? |  |  |  | 2.26 – 2.32 |
| Are all log books including testing, service, maintenance and repair records retained in the practice for at least 2 years? | Tick 1,2,3 or 4 1=all records available 2=most records available 3=some records available 4=no records available | 3.19 4.3, 4.15 |
| Cleaning: |  |  |  |  |
| Are disposable instrument trays used or if re-usable trays are used are they decontaminated and sterilized after each use? |  |  |  | 2.14 |
| Are any instruments (used or unused) left on trays at the end of each session decontaminated (washed and sterilized) before further use? |  |  |  | 2.4k |
| Are instruments that are not decontaminated immediately kept moist until they are decontaminated? | Tick 1,2,3 or 4 1=0 hours 2=1-3 hour 3=3-6 hours 4=6+hours | 2.15 3.5, 3.6 |
| Are instruments inspected under illuminated magnification device for cleanliness and condition following cleaning? | Tick 1,2,or 3 1=Always 2=Almost always 3=Never | 3.18, 3.49, 3.50, 3.51, 3.52 |
| Are hand pieces decontaminated in-between each patient in accordance to manufacturer’s instructions? |  |  |  | 2.10 note |
| Are separate canisters of lubricant used for unclean, cleaned and sterilized instruments? |  |  |  | 3.56 |
|  | YES | NO | N/A | Reference in HTM |
| Are those hand pieces that are manually cleaned/wiped, lubricated with oil before steam sterilization in accordance with manufacturer’s instructions? |  |  |  | 18.0 3.24, 3.55, 3.56 |
| Are those hand pieces decontaminated by an automated washer disinfector lubricated with oil before steam sterilization in accordance with manufacturer’s instructions? |  |  |  | 3.24 |
| Are those hand pieces decontaminated by an automated washer disinfector with a specific hand piece irrigation system, lubricated with oil before steam sterilization in accordance with manufacturer’s instructions? |  |  |  | 3.42, 3.24, 3.21 |
| Are those dental hand pieces washed by a specific hand piece washer device, lubricated with oil before steam sterilization in accordance with manufacturer’s instructions? |  |  |  | 3.22 |
| Are all dental instruments washed in a washer disinfector before steam sterilization? |  |  |  | 3.1, 3.2, 3.42, 4.3 |
| Where practices do not have a washer disinfector are all instruments cleaned (manually or using an ultrasonic cleaner) before steam sterilization? |  |  |  | 2.4(h) 3.33, 3.42 |
| Manual cleaning: |  |  |  |  |
| Are two sinks or two bowls in a single sink unit used for cleaning – one for washing and a separate one for rinsing? |  |  |  | 2.4h, 3.42, 16.1 |
| Are detergents used specifically formulated for purpose of cleaning instruments? |  |  |  | 16.3a |
| Is detergent used at specified concentration according to manufacturer’s guidance? |  |  |  | 16.3a |
| Is the temperature of water 45˚C or lower? |  |  |  | 16.3b |
| Where manufacturer’s instructions permit are instruments fully submerged when cleaned? |  |  |  | 16.3c |
| Are brushes used to clean instruments single use or washed after each use and replaced at manufacturer’s recommended interval or when damaged? |  |  |  | 16.3f |
| Validation and testing: |  |  |  |  |
| Are there contractual arrangements to ensure all steam sterilizers are routinely maintained and validated in accordance with HTM requirements or with manufacturer’s instructions? | Tick 1,2,3 or 4 1=Maintained and validated 2=Maintained 3=Validated but not maintained 4=Separate arrangements for maintenance and validation | 1.9, 3.11 11.1, 12.0 |
|  | YES | NO | N/A | Reference in HTM |
| Are daily, weekly, quarterly and annual inspection, testing and maintenance records available for steam sterilizers as described in section 12 of the HTM 01-05? | Tick 1,2,3 or 4 1=All records available 2=Most records available 3=Some records available 4=No records available | 12.0 |
| Is the steam sterilizer removed from service following an unsatisfactory test result until the fault is rectified? |  |  |  | 11.2 4.23 |
| Are there arrangements to ensure all ultrasonic cleaners are maintained and validated in accordance with HTM 01-05 requirements or with manufacturer’s instructions? |  |  |  | 14.0 15.6 |
| Are daily, weekly, quarterly and annual inspection, testing and maintenance records available for ultrasonic cleaners? | Tick 1,2,3 or 4 1=All records available 2=Most records available 3=Some records available 4=No records available |  |
| Are there contractual arrangements to ensure all automated washer disinfectors are routinely maintained and validated in line with manufacturer’s instructions? | Tick 1,2,3 or 4 1=Maintained and validated 2=Maintained 3=Validated but not maintained 4=Separate arrangements for maintenance and validation | 13.0 |
| Are daily, weekly, quarterly and annual validation and testing results recorded for automated washer disinfectors? | Tick 1,2,3 or 4 1=All records available 2=Most records available 3=Some records available 4=No records available | 13.0 |
| Ultrasonic Cleaners: |  |  |  |  |
| Are instruments placed in instrument basket or cassettes and fully immersed ensuring that all surfaces are in contact with the solution? |  |  |  | 3.30d-3.30e |
| Is the lid closed during cleaning cycles and whilst not in use to prevent contamination of the ultrasonic cleaning solution? |  |  |  | 3.30h |
| Is the water in the chamber emptied when visibly contaminated or otherwise at the end of every clinical session? |  |  |  | 3.30k |
| Where instruments are manually cleaned, are they rinsed after being ultrasonic cleaned and before sterilization? |  |  |  | 3.30m, 3.31 |
| Thermal Washer Disinfectors: |  |  |  |  |
| Are relevant staff aware of the instrument loading procedure i.e. spray arms are free to rotate, cannulated instruments are correctly loaded? |  |  |  | 3.17 |
| Are record cycle parameters recorded? |  |  |  |  |

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Sterilizers: |  |  |  |  |
| Is there a record made of the date, temperature and pressure achieved and satisfactory completion for each cycle? | Tick 1,2,3 or 4 1=All records available 2=Most records available 3=Some records available 4=No records available | 4.3, 4.14, 4.16 |
| Are steam sterilizers used if fault lights are displayed? |  |  |  | 4.23 |
| Are pre-wrapped instruments placed only in vacuum type sterilizers? |  |  |  | 4.11 |
| Is freshly distilled water, sterile water for irrigation or reverse osmosis (RO) water used in the sterilizer? |  |  |  | 4.13 |
| Are opened bottles of sterile or distilled water discarded at the end of each working day? |  |  |  | 17.6 |
| Is the reservoir drained and left clean and dry at the end of each day? |  |  |  | 4.13 |
| Decontamination environment: |  |  |  |  |
| Is there a zoned workflow from dirty to clean? |  |  |  | 5.3, 5.6, 5.7 |
| Are there separate, dedicated decontamination room/s which are restricted to those performing decontamination duties? |  |  |  | 1.10 |
| Are decontamination areas and work surfaces clean and uncluttered? |  |  |  | 5.6 |
| Is there adequate ventilation in the clean and dirty room/s to service W/D and sterilizer? |  |  |  | 6.42 |
| Where full mechanical ventilation is used is direction of air flowing from the clean area to dirty? |  |  |  | 6.44, 6.45 |
| Are there procedures in place for the safe transfer of instruments within the practice? |  |  |  | 2.26, 2.27 |
| Are instruments maintained in a moist condition between use and decontamination? |  |  |  | 2.15, 3.5, 3.6 |
| If transport containers are in use, are they lidded, clean, leak proof and in good working order? |  |  |  | 2.27 |
| Are transport containers cleaned, disinfected and dried following each use? |  |  |  | 2.28 |
| Are instruments processed in non vacuum (type N) sterilizer dried prior to packing using disposable non linting cloth? |  |  |  | 4.25 |

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have a system in place to ensure that storage of non-wrapped and wrapped instruments does not exceed 1 year for those instruments sterilized in non vacuum sterilizers (type N)? Or 1 year if sterilized in a validated type B vacuum sterilizer or in a cassette following sterilization in a validated type S sterilizer? |  |  |  | 4.31 |
| Is there a system in place to ensure that wrapped instruments are stored away from the clinical environment and used in strict rotation? |  |  |  | 1.10, 4.24 |
| For each instrument is there a system in place to identify storage time, including the date by which they should by used or reprocessed? |  |  |  | 4.24 |
| Are instruments stored in a dedicated secure, dry and cool environment? |  |  |  | 4.29 |

- Environmental design and cleaning

The dental environment is designed, maintained and cleaned appropriately to reduce the risk of cross infection

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have a policy and procedure for cleaning and maintaining the environment? |  |  |  | 2.6 6.54 |
| Have staff undertaking cleaning duties been fully trained to undertake such duties? |  |  |  | 6.55 |
| Is the overall appearance of the clinical and decontamination environment tidy and uncluttered? |  |  |  | 5.6 |
| Is the dental chair cleaned between each patient? |  |  |  | 6.46 |
| Is the dental chair free from rips or tears? |  |  |  | 6.62 |
| Are all surfaces, i.e. walls, floors, ceilings, fixtures and fittings and chairs free from damage and abrasion? |  |  |  | 6.39 |
| Are all work surface joints intact and seamless with no visible damage? |  |  |  | 6.46 |
| Are all surfaces, i.e. walls, floors, ceilings, fixtures and fittings and chairs free from dust and visible dirt? |  |  |  | 6.39 |
| Are the surfaces of accessible ventilation fittings/grills cleaned weekly? |  |  |  | 6.64 |
| Are all surfaces in clinical and decontamination areas impervious and easy to clean? |  |  |  | 6.64 |
| Are keyboard covers or “easy clean” waterproof keyboards used in clinical areas? |  |  |  | 6.66 |
| Are rooms where clinical practice takes place carpeted? |  |  |  | 6.46 |
| Do all floor coverings in clinical and decontamination areas have coved edges that are sealed and impervious to moisture? |  |  |  | 6.47 – 6.49 |
| Are soft toys available? |  |  |  | 6.73 |
| Are free standing or ceiling mounted fans used in clinical/decontamination areas? |  |  |  | 6.41 |
| Are records of cleaning maintained in accordance with HCAI Code of Practice? |  |  |  | 6.54 |
| Is cleaning equipment colour coded in accordance to the National Patient Safety Agency’s recommendations as detailed in HTM 01-05? |  |  |  | 6.53 |
| Is cleaning equipment stored in a non clinical area? |  |  |  | 6.60 |
| Where disposable single use covers are used, are they discarded after each patient contact? |  |  |  | 6.65 |
|  | YES | NO | N/A | Reference in HTM |
| Are the surfaces of equipment cleaned between each patient? E.g. work surfaces, dental chairs, curing lamps, delivery units, inspection handles and lights, spittoons, external surface of aspirator and X-ray heads? |  |  |  | 6.62 |
| Are all taps, drainage points, splashbacks, sinks, aspirators, drains, spittoons cleaned after every session with a surfactant/detergent? |  |  |  | 6.63 |
| Are floors, cupboards doors and accessible high level surfaces and floors cleaned daily? |  |  |  | 6.63 |
| Are floor coverings in clinical and decontamination areas impervious and easy to clean? |  |  |  | 6.46, 6.47 6.49 |
| Is there a designated area for the disposal of dirty water which is outside the kitchen, clinical and decontamination areas, for example toilet, drain, slop hopper (a device used for the disposal of liquid or solid waste) to reduce the risk of contamination of a public or staff toilet? |  |  |  |  |
| Does the practice have a local policy and procedure/s for spillage in accordance with C.O.S.H.H? |  |  |  | 2.4d, 2.6 |

4. Hand Hygiene

Hands will be decontaminated correctly and in a timely manner using a cleansing agent to reduce risk of cross infection

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have a local policy and procedure/s for hand hygiene? |  |  |  | 2.6 Appendix 2 |
| Is hand hygiene an integral part of staff induction? |  |  |  | 6.3 |
| Is hand hygiene training provided periodically throughout the year? |  |  |  | 1.25, 6.3 |
| Is hand hygiene carried out before and after every new patient contact? |  |  |  | Table A1 Appendix 2 |
| Is hand hygiene performed before donning and following the removal of gloves? |  |  |  | 6.5 Appendix 2 |
| Do all staff involved in any clinical and decontamination procedures have short nails that are clean and free from nail extensions and varnish? |  |  |  | 6.9, 6.24 Appendix 2 |
| Do all clinical and decontamination staff remove wrist watches, wrist jewellery, rings with stones during clinical and decontamination procedures? |  |  |  | 6.10, 6.23 |
| Are there laminated or wipeable posters promoting hand hygiene on display? |  |  |  | 6.13 |
| Is there a separate dedicated hand basin provided for hand hygiene in each surgery where clinical practice takes place? |  |  |  | 2.4g 6.11 |
| Is there a separate dedicated hand basin available in each room where the decontamination of equipment takes place? |  |  |  | 2.4u 6.11, 5.7 |
| Are wash hand basins free from equipment and other utility items? |  |  |  | 2.4g |
| Is bar soap available at wash hand basins? |  |  |  | Appendix 1 6.6 |
| Are hand hygiene facilities clean and intact (check sink taps, splash backs, soap and paper towel dispensers)? |  |  |  | 6.63 |
| Are there plugs and overflows on wash hand basins? |  |  |  | 6.11 |
| Does the water from the tap discharge away from the drain aperture? |  |  |  | 6.11 |
| Are elbow/wrist/foot operated, electronic mixers or thermostatically controlled taps available at all wash hand basins in clinical and decontamination areas? | Tick 1,2 or 3 1 = All 2 = Some 3 = None | 6.11 |
| Are nailbrushes present at wash hand basins |  |  |  | Appendix 2 |
| Is there good quality, mild liquid soap dispensed from single-use cartridge or containers available at each wash hand basin? |  |  |  | 6.6 Appendix 2 |
|  | YES | NO | N/A | Reference in HTM |
| Is skin disinfectant rub/gel available at the point of care? |  |  |  | Appendix 2 |
| Are good quality disposable absorbent paper towels used at all wash hand basins? |  |  |  | 6.7 Appendix 2 |
| Are hand-cream dispensers with disposable cartridges available for all clinical and decontamination staff? |  |  |  | 6.8 Appendix 2 |

5.  Management of Dental Medical Devices – equipment and dental instruments

Dental medical devices are operated, maintained, serviced and repaired to ensure adherence to patient safety and manufacturer’s instructions

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have an infection control policy that includes procedures for the use, maintenance, service and repair of all medical devices? |  |  |  | 1.20, 2.4a, 2.6, 2.7, 3.54 |
| Does the practice identify an individual with nominated responsibility and authority to ensure that all staff comply with the medical device procedure? |  |  |  | 2.4c |
| Has the practice carried out a risk assessment for legionella under the Health & Safety Commission’s ”Legionnaires’ disease – the control of legionella bacteria in water systems: Approved Code of Practice & Guidance” (also known as L8)? |  |  |  | 6.75-6.90, 19.0 |
| Has the practice a written scheme for prevention of legionella contamination in water pipes and other water lines? |  |  |  | 6.75, 19.2 |
| Are all new re-usable instruments decontaminated prior to use? |  |  |  | 2.6, 3.4 10.24 |
| Are contaminated medical devices decontaminated and inspected prior to inspection, maintenance and repair? |  |  |  | 3.54 |
| Are instruments sent for repair labelled to identify that they have been through the decontamination process? |  |  |  | 3.54 |
| Are single use instruments re-processed? |  |  |  | 2.17 |
| Are endodontic files and reamers re-used? |  |  |  | 2.21 |
| Dental radiography: |  |  |  |  |
| Are intra oral films, digital sensors and cassettes handled and stored safely in accordance with manufacturer’s instructions and to reduce cross infection? |  |  |  | 6.73 |
| Are film holders used in intra oral radiography subject to sterilization after every patient use in accordance with manufacturer’s instructions? |  |  |  | 6.72 |
| Impression material, prosthetic and orthodontic appliances: |  |  |  |  |
| Are impression materials, prosthetic and orthodontic appliances decontaminated in the surgery prior to despatch to laboratory in accordance with manufacturer’s instructions? |  |  |  | 7.0 |
| Are prosthetic and orthodontic appliances decontaminated before being placed in the patient’s mouth? |  |  |  | 7.1b |

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Other Medical Devices: |  |  |  |  |
| Are single-use items only used for single treatment episode and disposed of following use? |  |  |  | 2.17 |
| Are endodontic reamers and files treated as single use and disposed of following use? |  |  |  | 2.21 |
| Are difficult to clean instruments/devices (e.g. matrix bands, saliva ejectors, aspirator tips and three-in-one tips, etc,) identified as single-use? |  |  |  | 2.20 |
| Dental Unit Waterlines (DUWLs) |  |  |  |  |
| Are in-line filters cleaned/replaced as per manufacturer’s instructions? |  |  |  | 6.89, 6.90 |
| Is there an independent bottled water system used to dispense fresh distilled, reverse osmosis (RO) or sterile water to supply the DUWL? |  |  |  | 6.84 note |
| For dental surgical procedures involving irrigation, is a separate single use sterile water source used for irrigation? |  |  |  | 6.91 |
| Are the DUWLs drained down at the end of every working day? |  |  |  | 6.82 |
| Are self-contained water bottles (bottled water system) removed, flushed with distilled or clean RO water and left open to the air for drying on a daily basis and, if necessary overnight and in accordance with manufacturer’s guidance? |  |  |  | 6.83 |
| Where bottled water systems are not used is there a physical air gap separating dental unit waterlines from mains water systems (Type A)? |  |  |  | 6.84 note |
| Are DUWLs flushed for 2 minutes at start of each working day and for 20-30 seconds between every patient? |  |  |  | 6.85 |
| Are all DUWL and hand pieces fitted with anti-retraction valves? |  |  |  | 6.87 |
| Are DUWLs either disposable or purged using manufacturer’s recommended disinfectants? |  |  |  | 6.84 – 6.86 |
| Are DUWL filters changed according to the manufacturer’s guidelines? |  |  |  | 6.89 |
| Inhalation Sedation Machines (ISM): |  |  |  |  |
| Are ISM and accessories (including tubing, masks, nasal hood and nose pieces) used in accordance with manufacturer’s or supplier’s instructions? |  |  |  |  |
| Are ISM flow meters used and maintained in accordance with original equipment manufacturer’s or supplier’s instructions? |  |  |  |  |

6. Personal Protective Equipment

Personal protective equipment is available and is used appropriately to reduce the risk of cross infection

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have a policy and procedure/s for the use of personal protective equipment? |  |  |  | 2.6, 6.37 |
| Are staff trained in the use of personal protective equipment as part of the practice induction? |  |  |  | 6.14 |
| Are powder free CE marked gloves used in the practice? |  |  |  | 6.21 |
| Are alternatives to latex gloves available? |  |  |  | 6.19, 6.20 |
| Are all single-use PPE disposed of after each episode of patient care? |  |  |  | 6.22 |
| Is hand hygiene performed before donning and following the removal of gloves? |  |  |  | 6.5 Appendix 2 |
| Are clean, heavy duty household gloves available for domestic cleaning and decontamination procedures where necessary? |  |  |  | 6.24 |
| Are the heavy duty household gloves washed with detergent and hot water and left to dry after each use? |  |  |  | 6.24 |
| Are the heavy duty household gloves replaced weekly or more frequently if worn or torn? |  |  |  | 6.24 |
| Are disposable plastic aprons worn during all decontamination processes or clinical procedures where there is a risk that clothing/uniform may become contaminated? |  |  |  | 6.15 6.25 – 6.26 |
| Are single-use plastic aprons disposed of as clinical waste after each procedure? |  |  |  | 6.26 |
| Are plastic aprons, goggles, masks or face shields used for any clinical and decontamination procedures where there is a danger of splashes? | Tick 1,2,or 3 1=All 2=Most 3=None | 6.15 6.27 – 6.30 |
| Are all masks disposed of as clinical waste after each use? |  |  |  | 6.37 |
| Are all items of PPE stored in accordance with manufacturer’s instructions? |  |  |  | 6.15 |
| Are uniforms worn by all staff changed at the end of each day and when visibly contaminated? |  |  |  | 6.35 |
| Is eye protection for staff used during decontamination procedures cleaned after each session or sooner if visibly decontaminated? |  |  |  | 6.30 |
| Is eye protection provided for the patient and staff decontaminated after each episode of patient care? |  |  |  | 6.30 |

7. Waste

Waste is disposed of safely without the risk of contamination or injury in accordance with legislation

|  | YES | NO | N/A | Reference in HTM |
| --- | --- | --- | --- | --- |
| Does the practice have a policy and procedure/s for the management and disposal of waste? |  |  |  | 2.6 |
| Have all staff attended induction and ongoing training in the process of waste disposal? |  |  |  | 1.25 |
| Is there evidence that the waste contractor is a registered waste carrier? |  |  |  | Appendix 1 |
| Is the practice registered with Environment Agency if generating of 500kg per annum of hazardous waste? |  |  |  | Appendix 1 |
| Are all disposable PPE disposed of as clinical waste? |  |  |  | Appendix 1 6.26, 6.28, 6.37 |
| Are orange bags used for infectious Category B waste such as blooded swabs/blood contaminated gloves and teeth without amalgam fillings? |  |  |  | Appendix 1 |
| Are yellow/black bags used for offensive/hygiene waste such as non infectious recognisable healthcare waste, e.g. gowns, tissues, non contaminated gloves, x-ray film, etc, which are not contaminated with saliva, blood, medicines, chemicals or amalgams |  |  |  | Appendix 1 |
| Are black/clear bags used for domestic waste including paper towels? |  |  |  | Figure A1, Appendix 1 |
| Are bins foot operated or sensor controlled, lidded and in good working order? | Tick 1,2,or 3 1 = All 2 = Some 3 = None | Appendix 2 |
| Are local anaesthetic cartridges and other Prescription Only Medicines (POMs) disposed of in yellow containers with a yellow lid that conforms to BS 7320 (1990)/UN 3291? |  |  |  | Appendix 1 |
| Are clinical waste sacks securely tied and sharps containers locked before disposal? |  |  |  | Appendix 1 |
| Are all clinical waste bags and sharps containers labelled before disposal? |  |  |  | Appendix 1 |
| Is waste awaiting collection stored in a safe and secure location away from the public within the practice premises? |  |  |  | Appendix 1 |
| Are all clinical waste bags fully described using the appropriate European Waste Catalogue (EWC) Codes as listed in HTM 01-05? |  |  |  | Appendix 1 |
| Are all consignment notes for all hazardous waste retained for at least 3 years? |  |  |  | Appendix 1 |
| Has the practice been assured that a “duty of care” audit has been undertaken and recorded from producer to final disposal? |  |  |  | Appendix 1 |
| Is there evidence the practice is segregating waste in accordance with HTM 01-05? |  |  |  | Appendix 1 |